## Sheet1
ECG MACHINE
SR # | SPECIFICATIONS
1 | Twelve Channel ECG on at least 12 inches touch screen display
2 | Automatic Operation
3 | Variable gain: 1/2, 1, 2 cm/mV
4 | Thermal recorder for printing out Twelve channels simultaneously and should have the capability to connect with external thermal printers.
5 | Interpretation software
6 | Recording Trace speed: 10, 25 and 50 mm/sec,
7 | Muscle artifact and AC (50Hz) interference filters
8 | Defibrillator protection
9 | Built-in AC operation & battery backup minimum 5 hours
10 | Paper size: A4/210mm
11 | Built-in AC interference, noise filter and baseline drift control.
12 | Capability to interface with LAN/WLAN for data transfer
13 | Accessories: Complete with standard accessories, including separate patient cables for Adult, Pediatric & Neonatal use with re-usable electrodes(procuring agency will specify the type of cable needed) Optional (If any): Mobile Cart (Local/Imported) No. of Electrodes